Resistive Transduction
Suresh Devasahayam
Department of Bioengineering
Christian Medical College, Vellore

---
Lecture - Outline
- Variable Resistance Potentiometer as Transducer
- Power and Error in Potentiometer measurement
- Measuring Small Resistance Changes
  - Bridge Measurement
  - Single Active Arm Bridge
- Temperature compensation
  - Two and Four Active Arm bridge

---
Displacement measurement with a potentiometer

---
Simple case
V_out = V_{i} over R_P cdot x R_P
V_out over V_{i} = x

---
Simple case:
- If RL>>RP :
- Current drawn by Rp :
- Power dissipated:
- To minimize power dissipated in the potentiometer, the value of RP should be large.
V_out = V_i over R_P cdot x R_P
V_out over V_i = x
I_P = V_i over R_P
Power_P = V_i ^2 over R_P

---
Considering the value of the measuring instrument
- RP is large :
- RL and RP are large:
- Measurement error:
V_out =I_{R_L} cdot R_L= V_i over { left((1-x)R_P + {xR_P R_L}over {xR_P+R_L} right) } cdot left( xR_P over {xR_P+R_L} right) cdot R_L
V_out over V_i = {xR_L} over {xR_P(1-x)+R_L}
V_out over V_%in = {x %alpha} over { x(1-x)+%alpha}
%alpha = R_L over R_P
Error={Actual value-Desired value} over {Desired value} = { left( {x %alpha} over {x(1-x)+%alpha} right) -x} over x = {-x(1-x)} over {%alpha +x(1-x)}

---
Bridge Measurement

---
Measuring small resistance change
V_o = V_a - V_b = V_i left ( Z_4 over {Z_1 + Z_4} - Z_3 over {Z_2+Z_3} right )

---
Measuring the bridge output
- Thevenin equivalent of the bridge
- Output impedance or source impedance =
  OC voltage/SC current
Z_s = {Z_1 Z_2} over {Z_1 + Z_2} + {Z_3 Z_4} over {Z_3+Z_4}

---
One Active Arm Bridge
A resistive transducer’s resistance varies around a “no signal” value of “R”
Z_4 = R +- %DELTA R = R(1 +- %DELTA),~~ Z_1=Z_2=Z_3=R
V_o=V_i %DELTA over {4+2%DELTA} approx V_i %DELTA over 4 ~,~~ Z_s=R {(4+3%DELTA)} over {(4+2%DELTA)} approx R

---
Two active arm bridge
- Resistance also has a temperature dependence
- If transduction can yield complementary resistance change, then we can use a two active arm bridge:
Z_4=R(1+%DELTA_s+%DELTA_T),~~ Z_3=R(1-%DELTA_s+%DELTA_T),~~ Z_1=Z_2=R
V_o=V_i left LBRACE {R (1+%DELTA)} over {R(2+%DELTA)} - {R (1-%DELTA)} over {R(2-%DELTA)} right RBRACE
V_o=V_i left LBRACE {(1+%DELTA)(2-%DELTA)-(1-%DELTA)(2+%DELTA)} over {(4-%DELTA^2)} right RBRACE =V_i {2 %DELTA} over {4-%DELTA^2} approx V_i %DELTA over {2}~
V_o approx V_i %DELTA over 2 ~,~~ Z_s=R(1- %DELTA ^2 over 2)