Resistive Displacement & Force Sensors
Strain Gauges
Suresh Devasahayam
Department of Bioengineering
Christian Medical College, Vellore

---
Lecture - Outline
- Strain gauges
  - Load cells
  - Force measurement
  - Pressure measurement
- Thermistors
  - Temperature measurement
  - Self-heating

---
Strain Gauge
- Resistance
- Taking log
- Differentiating
- Fractional change of resistance
- Can calculate for a cylinder, A=πr2
- Poisson’s ratio σ is 0.5 for incompressible solids (constant volume), 0.3 for metals
R=%rho l over A
ln R = ln %rho + ln l - ln A
{d ln R} over dR =1 over R, Thus: d{ln R} = dR over R
dR over R = {d %rho} over %rho + {d`l} over l - dA over A
{%DELTA R} over R = {%DELTA l} over l - {%DELTA A} over A + {%DELTA %rho} over %rho
{%DELTA r} over r = - %sigma {%DELTA l} over l
{%DELTA A} over A approx 2 {%DELTA r} over r
{%DELTA R} over R = {%DELTA l} over l + 2 %sigma {%DELTA l} over l + {%DELTA %rho} over %rho

---
Strain Gauge (contd)
- Temperature dependence of resistivity
- Temperature dependence of length
%rho = %rho_o [1+%alpha_R(T-T_o)]=%rho_o + %DELTA %rho_T
l = l_o [1+%alpha_l(T-T_o)]
{%DELTA R} over R = {%DELTA l} over l (1 + 2 %sigma ) + {%DELTA %rho_T} over %rho = G {%DELTA l} over l + %alpha_R %DELTA T

---
Strain Gauge Examples
Material	Temperature Coeff / αE	Maximum strain / Δl/l	Modulus of Elasticity / Y	Gauge factor / G
Platinum	0.0038	0.0023	1.5 x 1010	6
Nichrome	0.0004	0.0036	19 x 1010	2.5
Silicon	0.007	0.0033	19 x 1010	170

---
Load Cell

---
Load Cell – cantilever beam

---
Bending of the cantilever beam
- Deflection of beam
- Horizontal distance:
- Similar Sectors:
- Substiting for r:
- Strain in surface fibres:
%delta = {W l^3} over {3 Y I_xx}
~~where ~~ I_xx=1 over {12} bh^3
l^2 - %delta ^2 = r^2 - (r-%delta)^2
{%DELTA l} over l = { h/2} over r
{%DELTA l} over l = {h %delta} over {l^2}
{%DELTA l} over l = {h %delta} over { l^2} = {4W} over {Ybh^2}

---
Mechanical Equivalent
F(t) = K x(t) + B {dx(t)} over dt +M {d^2 x(t)} over {dt^2}
X(s) = {F(s)} over { K+Bs+Ms^2}
X_o(s) = {A F(s)} over {s^2 over {%omega_n^2} + 2 %zeta over %omega_n s + 1}

---
Pressure Transducer

---
Pressure Measurement

---
Equivalent of catheter fluid system
L_c={%rho l} over A
C_c = {Ald} over {Eh}
R_c = {8 %pi %eta l} over {A^2}
C_d = 1 over {E_d}
C_b = {%DELTA V_b} over {%DELTA P_b}

---
Thermistors –
negative temperature coefficient sensors

---
Thermistors

---
Thermistors
- Thermistor characteristics
- Sensitivity:
- Non-linear sensitivity:
R_T = R_o e^{%beta (1 over T - 1 over T_o)}
ln R_T - ln R_o = %beta over T - %beta over T_o
1 over R_T {dR_T} over {dT} - 0 = %beta {-1 over {T^2} -0}
1 over R_T {dR_T} over {dT} = -%beta over T^2
{%DELTA R_T} over R_T = {%DELTA T} left ( {-%beta over T^2} right )

---
Voltage-Current characteristics

---
Thermistor self-heating
- Power dissipation (due to electric current) results in heating of thermistor
- As temperature changes, resistance decreases
- Equilibrium is reached when the electrical power is equal to the thermal power dissipated
P=CM { dT } over {dt} +D( T-T_a)
where~C=sp.`heat, ~~M=mass, ~D=dissipation~constant
T(t)=T_a+P over D [ 1-e^{D over {CM} t} ]
steady~state~~: ~~D=P over { T-T_a }

---
Application of Thermistor in Self-Heating
- Measurement of fluid flow – changing dissipation constant
- Mass flow meters
  - gas and liquid flow sensors