Voice Coil Example
Suresh Devasahayam
Department of Bioengineering
Christian Medical College, Vellore

---
Lecture - Outline
- Voice Coil
  - Ring magnet
  - Centre pole piece
  - Air gap
  - Coil in air gap
  - Spring to position coil

---
Voice coil - actuator

---
Ring magnet
Ferrite ring
Length=20mm
outer diameter=20mm, inner diameter=15mm.
Place it in a coil of wire (e.g.,20 turns; current 1A)
Ferrite => Hc=60 Amp-turn/m, Br = 0.15 Weber/sq.m
Ferrite ring is saturated and fully magnetized.

---
Permanent Magnet
R_m = H_c over B_m l_m over A_m = 400 left ( {0.02} over {137.5x10^-6} right ) = 58.2 x 10^3 Amp-turns/Weber

---
Magnetic circuit
Iron Pole piece: length = 20mm, diameter = 13mm.
Fe permeability:
%mu=6 x 10^-3
R_p = 1 over %mu l_p over A_p = 10^3 over 6 0.02 over { 530 x 10^-6 } = 6.3 x 10^3
R_g = 1 over %mu_o l_g over A_g = 1 over { 4 %pi 10^-7 } {2x10^-3} over { 880 * 10^-6 }=1800x10^3
B_g = H_c over { R_m + R_p + R_g } R_g = H_c x 0.965 = 58 Wb/m^2

---
Coil
Coil: 40 gauge (diameter 0.08 mm), 125 turns. Length of coil=10mm.
Diameter of coil = 14 x 10-3 m. Length of wire=1.75 m
Resistance of coil = 5.8 Ω
Inductance (iron core inductor):
L = %mu N^2 A over l= 6 x 10^-3 x 15625 x {154 x 10^-6} over {10 x 10^-3} = 1.44 H

---
Static case
Excitation voltage = 6V
Current = 1.03 A
Force:
F=BIl= 105N

---
Dynamic case
Excitation voltage is time varying: V(t)
V( t ) = R~I( t )+L { dI(t) } over { dt }
I( s ) = { V(s) } over { R+Ls } = {V( s )} over {Z_e( s )}

---
Mechanical Aspect
- When coil is moving, EMF due to moving conductor in magnetic field:
  - Back EMF
  - Velocity determined by mechanics
V_b (s)=B_g l_c u(s)
V(s) = V_e(s) + V_b (s) newline ~~~~~~~~~~= {I(s)} {Z_e(s)} + B_g l_c u(s)
F(t) = m dot u(t) + k_b u(t) + k_1 int u(t) dt newline u(s) = F(s)/[ms+k_b +k_1 /s] newline u(s) = {B_g l_c I(s)} over{ms+k_b +k_1 /s}

---
Equivalent Electrical Impedance due to mechanical components
{V_b(s)} over {I(s)}~=~-B_g ^2`l_c ^2`{Z_m(s)}

---
Using a transformer representation
V_2 = k V_1 ~,~~~~~I_2=I_1 over k ~,~~~~V_2 over I_2 = k^2 V_1 over I_1 ~,~~~Z_L=V_2 over I_2 ~,~~~V_1 over I_1 = Z_L over k^2

---
Servo Motor

---
Application of Servo motor in Haptics

---
Haptics – Feeling in Medical Devices

---
End of Lecture