Piezoelectricity
Suresh Devasahayam
Department of Bioengineering
Christian Medical College, Vellore

---
Lecture - Outline
- Piezoresistivity and piezoelectricity
- Electrical characterisitics
- Applications

---
Resistivity – Ohm’s law
- Isotropic material
E over J = %rho
{V/l} over {I/A}=%rho~, V over I = %rho l over A=R

---
Resistivity – non-isotropic material
- General non-isotropic material
  - e.g., biological tissue
E_1 = %rho_11 J_1 + %rho_12 J_2 + %rho_13 J_3 newline E_2 = %rho_21 J_1 + %rho_22 J_2 + %rho_23 J_3 newline E_3 = %rho_31 J_1 + %rho_32 J_2 + %rho_33 J_3
E_i = sum from{j=1,2,3} %rho_{ij} J_j

---
Piezoresistivity
- In material like silicon there is a mechanical stress related component of resistivity
- If there is stress (force/area) in direction k and a piezoresistive coefficient corresponding to normal and shear stresses, the effective electric field is:
E_i ~ = sum from{j=1,2,3} %rho_{ij} J_j ~+~ sum from {j=1,2,3}` sum from {k=1,...,6} %pi_{ijk}J_i %phi_k

---
Polarization of charge in a dielectric and Mechanical stress in solids
- Charge per unit area depends on the permittivity and the electric field
- Capacitance
- Stress in solids
Q over A = %epsilon E
C=Q over V = %epsilon A over l
{%phi}={%DELTA f} over {A} = Y {%DELTA l} over l

---
Piezoelectricity
- Stress causes additional charge separation
- Electric field causes additional strain
Q over A = %epsilon E + d_p %phi
{%DELTA l} over l = %phi over Y - d_p E

---
For general non-isotropic piezoelectric material
- Charge separation is affected by directional stress
- Strain is affected by directional electric field
Q over A = P_i = sum from {j=1...3} %epsilon_ij E_j + sum from {m=1...6} d_{i m} %phi_m
{%DELTA l} over l = sum from{m=1...6} %phi_i cdot {1 over Y_{i m}} + sum from {j=1...3} d_ij E_i

---
Electrical Equivalent of Piezoelectric crystal
%DELTA Q(t) = d_p cdot F(t) ~,~~~~ C_o = %epsilon {bl} over h

---
Example
Ammonium Dihydrogen Phosphate (ADP): dp-36 = 24 x 10-12 Coulombs/Newton
(i) dimensions: l=10mm, b=10mm, h=10mm
Applied voltage of 1000V.
Applied stress is zero, and the electric potential is applied in z-direction (direction 3).Deflection is a shear about the z-axis (direction 6)
The resulting strain is:
(ii) dimensions: l=1mm, b=10mm, h=10mm
Strain = %phi over Y + d_p E
(i) strain =24 cdot 10^{-12} 1000 over {10^-2} newline ~= 24 cdot 10^{-7} radians
(ii) strain = 24 cdot 10^{-6} radians

---
Mechanical resonance
- Beam with mass, elasticity and damping
- Exchange of energy between mass and elastic spring
- Velocity of sound in a solid
- Oscillation of a beam with a half-wave
  (both ends fixed)
{X (s)} over {F (s)}= A over {ms^2 +B s + K}
{X(s)} over {F(s)}=A over { s^2 over %omega_c ^2 + {2 %zeta} over %omega_c s + 1}
u~=~ sqrt{~ Y over %rho}
l= %lambda over 2 = u over {2 f_r}~, ~~~ f_r = u over {2 l}

---
Example 1
- Quartz:
  - Speed of sound=5800m/s
- Beam of length:
  - l=1mm for full wavelength
  - L=0.5mm for half wave
  - L=0.25 for quarter wave
- Frequency of oscillation
  - 5.8MHz

---
Output Impedance of a piezoelectric crystal
- Mechanical impedance of the crystal
- If the end has zero force, then the mechanical impedance is zero (short circuit)
- If the end has zero velocity, then th mechanical impedance is infinity (open circuit)
Z_o = Y over u = sqrt {%rho Y} = %rho u
Stress = k Z_o u
Z_o = Stress over {(strain) cdot u}
Z_o = %rho u

---
Reflection
Reflectivity
R={ Z_2 - Z_1 } over { Z_2+Z_1 }

---
Applications of Piezoelectric crystals
- Ultrasound
  - Transit time measurement of distance
  - Doppler measurement of motion

---
End of Lecture