Electrical Safety
Suresh Devasahayam
Department of Bioengineering
Christian Medical College, Vellore

---
Introduction
Electrical Hazard
Electrical excitation of tissue – current, voltage, frequency
Electrical power transfer – burns
Electrical Instruments
Grounding instruments
Patient protection – electrical isolation
Isolation transformers
Optical, magnetic and capacitive isolation

---
Safety
- Electrical Safety in Hospitals
- The Mains Power Distribution
  - Electrical power distribution – three phase and single phase systems
  - Ground – Sub-station ground, local ground, neutral
- Leakage Current Detection and Protection Devices
- Ground Fault Circuit Interrupter (GFCI) or Residual Current Circuit Breaker (RCCB)

---
Danger to Humans – Biological reasons
Electrical Danger to Humans
Electrical Stimulation of nerves and muscles
Skeletal Muscle contraction
Cardiac Muscle fibrillation

---
3 Phase Mains Power

---
Mains Power
- In India the nominal phase voltage is 240V @ 50Hz. (Voltage between any two phases is 415V)
- The Neutral is Ground-connected at the sub-station.
- The Neutral is the current return line for a single phase line.
- In a building, the individual power points are connected to Ground.
- All instrument bodies are connected to Ground.
- The Ground and Neutral should be at the same electrical potential.

---
Residual Current Testing
If there is no leakage to ground, I1=I2.
Residual Current Devices can be used to detect ground line leakage, and imbalance in 3-phase circuits (here the neutral may not always be balanced, even without fault- therefore, neutral should also be monitored).
MCB, ELCB, RCCB

---
Leakage Currents

---
Danger with Multiple Grounds

---
Danger to Humans – Kinds of Danger
Rules of Thumb
For electrical current entering from the surface of the skin
Less than 1mA: Microshock
1mA: Threshold of perception
5mA: Maximum harmless current
10mA: Let-go current
100mA: Ventricular fibrillation
5A: Burns
For electrical current directly into the heart (e.g., via a cardiac catheter)
Safe current is about 10μA