Department of Bioengineering, CMC-Vellore

Transducers and Instrumentation for Physiological Measurement
Mid-Term Test
1st March, 2019
1 hour

The following table gives the resistance of a thermistor measured at 3 different temperatures. Calculate the value of the thermistor constant β.
T (oC)	20	30	40
RT (Ω)	2492	1611	1060

The Resistance vesus Temperature relation is approximated by the expression: RT = 1611-55.1(T-30) where the temperature is in Celsius. This thermistor is to be used in the range 35o to 42oC. What will be the maximum error in this range, and at what temperature?
1. The resistance variation of a strain gauge gauge is given by: {%DELTA R} over R = G {%DELTA l} over l `+` %alpha (T-T_o)
  The gauge factor of strain gauges made of Nichrome and Silicon are 2.5 and 170 respectively, while the temperature coefficients of resistivity are 0.0009 and 0.007, with 20oC being the reference temperature.
  (a) Derive the expression for the output of a Wheatstone bridge for a load cell with (i) one strain gauge, and (ii) two strain gauges. (b) If 1% strain is present and 5oC temperature variation exists, calculate the following. Use an excitation voltage of 5V.
	Output due to 1% strain	Output due to 5oC	Total Output	Apparent strain
One active arm bridge
Two active arm bridge

1. We have two piezoelectric crystals: (a) quartz crystal of dimension x=5mm, y=5mm, z=2mm, and (b) barium titanate ceramic of dimension x=2mm, y=2mm, z=1mm. If a force of 100N is applied to each of them in the x direction for quartz and in the z direction for the ceramic crystal, what is the voltage in the preferred direction ? The value of the piezoelectric coefficient is: for Quartz, and dp(33)=-for Barium Titanate Ceramic. Permittivity of free space, 8.8542×10−12 F⋅m−1
	Relative permittivity (Dielectric constant)	Piezoelectric constant
Quartz (dp-11)	4.5	dp(11)=2.3x10-12 Coulombs/N
Barium Titanate Ceramic ()	1800	dp(33) =78x10-12 Coulombs/N