Department of Bioengineering, CMC-Vellore

Transducers and Instrumentation for Physiological Measurement
Mid-Term Test
1st March, 2019
1 hour

The following table gives the resistance of a thermistor measured at 3 different temperatures. Calculate the value of the thermistor constant β.
T (oC)	20	30	40
RT (Ω)	2492	1611	1060

The Resistance vesus Temperature relation is approximated by the expression: RT = 1611-55.1(T-30) where the temperature is in Celsius. This thermistor is to be used in the range 35o to 42oC. What will be the maximum error in this range, and at what temperature?
1. The resistance variation of a strain gauge gauge is given by: {%DELTA R} over R = G {%DELTA l} over l `+` %alpha (T-T_o)
  The gauge factor of strain gauges made of Nichrome and Silicon are 2.5 and 170 respectively, while the temperature coefficients of resistivity are 0.0009 and 0.007, with 20oC being the reference temperature.
  (a) Derive the expression for the output of a Wheatstone bridge for a load cell with (i) one strain gauge, and (ii) two strain gauges. (b) If 0.1% strain is present and 5oC temperature variation exists, calculate the following. Use an excitation voltage of 5V and R=100 Ohms.
	Output due to 0.1% strain	Output due to 5oC	Total Output	Apparent strain
One active arm bridge
Two active arm bridge

1. We have two piezoelectric crystals: (a) quartz crystal of dimension x=5mm, y=5mm, z=2mm, and (b) barium titanate ceramic of dimension x=2mm, y=2mm, z=1mm. If a force of 100N is applied to each of them in the x direction for quartz and in the z direction for the ceramic crystal, what is the voltage in the preferred direction ? Permittivity of free space, 8.8542×10−12 F⋅m−1
	Relative permittivity (Dielectric constant)	Piezoelectric constant
Quartz	4.5	dp(11)=2.3x10-12 Coulombs/N
Barium Titanate Ceramic	1800	dp(33) =78x10-12 Coulombs/N

Solutions

R_T=R_o e^{%beta (1 over T - 1 over T_o)} ,~~~~~~~~~ ln left [ R_T over R_o right ] = %beta (1 over T - 1 over T_o) ,~~~~~~~~~ R_T over 2492 = %beta (1 over T - 1 over 293)
Using To=20oC=293K, ln(R_T over 2492) = %beta (1 over T - 1 over 293)
and T=30oC=303K, -0.4362 = %beta (1 over 303 - 1 over 293)= %beta (-0.00011264) β=3872

  The actual resistance curve and the linear approximation curve plotted against temperature are shown here. The two graphs can be drawn quickly using the numbers given. The local maximum of the error has to be found, not the global maximum. The maximum error in the range 35-40 occurs around 35oC. In the range 35-42, the maximum is at 42oC.

Let the resistance variation due to strain and temperature be denoted by ΔRs and ΔRT respectively.
Strain component: {%DELTA _S} = G {%DELTA l} over l , Temperature component: {%DELTA _T} = %alpha ~ {%DELTA T}
Output of bridge with one strain gauge R±ΔRs±ΔRT, and three equal resistors, R is:
V_o = V_i left lbrace {+ %DELTA _s + %DELTA _T} over {4+%DELTA _s+%DELTA _T} right rbrace approx V_i left lbrace {+- %DELTA _s +- %DELTA _T} over {4} right rbrace
Two strain gauges are placed on a beam to experience strain in opposite directions so that the resistances are R+ΔRs+ΔRT and R-ΔRs+ΔRT respectively. A bridge made with these two gauges and two equal resistors, R will have output:
V_o = V_i left lbrace {2 %DELTA _s} over { (2+%DELTA_T)^2 - %DELTA_s ^2} right rbrace approx V_i left lbrace {2%DELTA_s} over {4} right rbrace
Therefore, in the two strain gauge configuration with the gauges experiencing opposite strains, the effect of temperature is cancelled.

Nichrome ( G=2.5, %alpha=0.0009 , Ve=5V)
	Output due to 0.1% strain	Output due to 5oC	Total Output	Apparent strain
One active arm bridge	3.125 mV	5.625 mV	8.75 mV	0.28%
Two active arm bridge	6.25 mV	0	6.25mV	0.1%

Silicon ( G=170, %alpha=0.007 , Ve=5V)
	Output due to 0.1% strain	Output due to 5oC	Total Output	Apparent strain
One active arm bridge	0.2125V	0.04375 V	0.25625 V	0.12%
Two active arm bridge	0.425 V	0	0.425V	0.1%

3. Q over A = sum from {j=1...3} %epsilon_ij E_j + sum from {m=1...6} d_{i m} %phi_m , Ej=0, Q over A = d_p F over A ~,~~~ Q=d_p F

V=Q over C,~~~~~C=%epsilon A over l

Quartz:
Q=2.3 x 10^-12 x 100=230x10^-12
C=4.5x8.8542x10^-12x(5x2)x10^-6 / 5x10^-3=79.6878x10^-15~,~~~V=2886 V

Barium Titanate Ceramic:
Q=78 x 10^-12 x 100,~~~Q=7800x 10^-12
C=1800x8.8542x10^-12x4x10^-6 / 1x10^-3=63750x10^-15~,~~~V=122 V